## Slide 1: MS Lifecam HD 3000 Determining FoV and specs
2022 1226
Cory Hill, Mentor, FRC 2429

## Slide 2: Camera calibration background
https://www.skiprobotics.com/articles/2021-08-13-camera-calibration/
https://ksimek.github.io/2013/08/13/intrinsic/

## Slide 3: Understanding the native cropping
MS Lifecam HD 3000

## Slide 4: Same scene overlay of 1280x720 (red outline) and 640x480 (added alpha, blue outline)
Will have different FoV in x and y.
Note how sensor crops x cols but GAINS y rows when going to 640x480

## Slide 5: Calculating the nominal intrinsics
MS Lifecam HD 3000

## Slide 6: 1280x720 FoV - 15.0” distance, 1” squares
10” in x = 738 pixels, so 73.8 per pixel. 1280/73.6 = 17.39” across
6” in y = 443 pixels, so 73.8 inches per. 720/73.6 = 9.78” high
Not accounting for distortion (need a bigger target/ get closer for that)
Fov_x = 2*arctan(17.39/2 / 15) = 60.2°
Fov_y = 2*arctan(9.78/2 / 15) = 36.1°
Diagonal Fov likely 2*arctan(1468.6/73.6/2/ 15) = 67.3 ° (within 1.2° of spec)

## Slide 7: 640x480 FoV - 15.0” distance, 1” squares
10” in x = 448 pixels, so 44.8 per pixel. 640/44.8 = 14.28” across
7” in y = 312 pixels, so 44.6 inches per. 480/44.6 = 10.76” high
Not accounting for distortion (need a bigger target / get closer for that)
Fov_x = 2*arctan(14.28 /2 / 15) = 50.1°
Fov_y = 2*arctan(10.76 /2 / 15) = 39.4°
Diagonal Fov likely 2*arctan(800/44.7/2/ 15) = 61.6 °

## Slide 8: Camera models – estimating the intrinsics
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
fx_predicted = 1104.1
fy_predicted = 1104.7
cx_predicted = 1280/2 = 640
cy_predicted = 720/2 = 360
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
fx_predicted = 684.7
fy_predicted = 670.2
cx_predicted = 640/2 = 320
cy_predicted = 480/2 = 240
fx_predicted = 0.5 * w / np.tan(0.5 * np.deg2rad(w_fov))
fy_predicted = 0.5 * h / np.tan(0.5 * np.deg2rad(h_fov))
[ 1104.1, 1104.7, 640, 360 ]
[684.7, 670.2, 320, 240 ]
We can feed these to apriltag.detector() for a generic Lifecam HD 3000
Distances will be quite close, but center offsets will not be exact

## Slide 9: Comparison of calibration techniques
MS Lifecam HD 3000

## Slide 10: Calibration techniques: 3DF Zephyr Free
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
fx_predicted = 1104.1
fy_predicted = 1104.7
cx_predicted = 1280/2 = 640
cy_predicted = 720/2 = 360
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
fx_predicted = 684.7
fy_predicted = 670.2
cx_predicted = 640/2 = 320
cy_predicted = 480/2 = 240
Nominal : [ 1104.1, 1104.7, 640, 360 ]
3DF: [1142.9, 1142.9, 650.9, 354.7]
evens [1144.5, 1144.5, 654.3, 364.6]
odds [1152.9, 1152.9, 649.4, 352.3]
Just the 4 biggest images: [1121.9 1121.9, 651.3, 360.4]
Nominal: [684.7, 670.2, 320, 240 ]
3DF: [681.6, 681.6, 322.6, 243.8]
Result with subsets of cal images:
Just evens [652.3, 652.3, 325.2. 237.3]
Just odds [692.4, 692.4, 210.2, 244.6]
Need to see which images cause the most noise, but 3DF Zephyr is pretty stable. Their cal image is shown on right. Cal wants a minimum of 3 images. More images seems to get to a more accurate result (as it should) – see later slides

## Slide 11: Calibration techniques: calibdb.net
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
Nominal : [ 1104.1, 1104.7, 640, 360 ]
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
Nominal: [684.7, 670.2, 320, 240 ]
Calibdb: [1120.5, 1119.7, 640.1, 353.7]
It’s cropping incorrectly from the 1280 data
Not sure how to fix that
Calibdb: [747.0, 746.4, 320.5, 235.8 ]
Has its own web interface at https://calibdb.net/ and custom target. Can’t get a good 640x480 cal out of calibdB – it can’t predict the way we’re cropping in slide 3

## Slide 12: Calibration techniques: ros2 cameracalibrator
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
Nominal : [ 1104.1, 1104.7, 640, 360 ]
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
Nominal: [684.7, 670.2, 320, 240 ]
ros2: [1080.8, 1076.8, 643.2, 351.3]
ros2:
first try [680.3, 685.8, 306.5, 244.6 ]
second try: [ 672.5, 672.5, 314.1, 252.2]
This one is quite nice – lets you know if you have enough data for a good cal. Takes a while to calculate though. However, gives the right matrices to feed to ROS’ aprilcal.
Had to set up ros2 on a pi and run everything.
Still has a some variation from cal to cal.

## Slide 13: Calibration techniques: opencv
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
Nominal : [ 1104.1, 1104.7, 640, 360 ]
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
Nominal: [684.7, 670.2, 320, 240 ]
:
:
I built the standard opencv demo pipeline on this, but the default method seems to be terrible – super unstable.
You can’t really get configuration wrong, because the only input is the images themselves
I can’t get anywhere close to the nominal prediction, regardless of image placement and tilt/skew.
Huge variation when leaving out one image
Tried with sets of 9 like most of what you see on the internet, then went up to 45 with 9 each of flat, tilt left, tilt right, tilt forward, tilt back
Nothing got me close to nominal
So I started changing the parameters
  If you supply a good guess and lock the aspect ratio, it gets a lot better

## Slide 14: Problems with calibration techniques
MS Lifecam HD 3000

## Slide 15: Noise and error with OpenCV
Can take a lot of pictures, with the advice often given online
  Same distance, get all of the viewing area, get tilted images in both x and y
  Check all the images, saw that they all recognize
Get truly wrong numbers from this method most of the time
Seems to benefit from a ton of images if the image is relatively small (like in the subset shown)
Have to get just the right set of images, otherwise fx and fy are really off

## Slide 16: Noise and error with OpenCV - continued
Take a large set and drop images and see how stable we are and get a sense of the sensitivity to one image
Different sets give wildly different fx and fy (and they aren’t the same, so the convention with OpenCV is possibly different from others)
You don’t really get ‘bad outliers’ – they’re usually the corner images that are keeping you closer to reality
Example numbers
  fx mean: 740.6 median: 665.3 std: 429.6
  fy mean: 1184.5 median: 1137.5 std: 127.2
  cx mean: 327.4 median: 321.6 std: 29.4
  cy mean: 194.6 median: 186.3 std: 18.3
  err mean: 57.9 median: 57.9 std: 0.11
The error is high, compared to what I see from good cal examples
Changing distances and going simple still gives results that are just plain wrong, so there must be a set of parameters that has to be changed to get this right

## Slide 17: Checking the intrinsics – fx and fy
MS Lifecam HD 3000

## Slide 18: Comparison of techniques with experiment – 640x480
h, w = 480, 640
w_fov = 50.1
h_fov = 39.4
Nominal: [684.7, 670.2, 320, 240 ]
Measured distance in z with Bosch laser = 4.00m. Calculated from AprilTag detector with cals below:
Summary of results
ID: 9 hamming: 2 decision_margin: 0.26
ID: 13 hamming: 2 decision_margin: 0.87
ID: 26 hamming: 0 decision_margin: 37.51

## Slide 19: Comparison of techniques with experiment – 640x480
Measured distance in z with Bosch laser = 4.00m. Calculated from AprilTag detector with cals below:
Summary of results
ID: 3 hamming: 2 decision_margin: 0.76
ID: 12 hamming: 2 decision_margin: 0.14
ID: 12 hamming: 1 decision_margin: 0.04
ID: 25 hamming: 2 decision_margin: 1.06
ID: 26 hamming: 0 decision_margin: 24.18
h, w = 720, 1280
w_fov = 60.2
h_fov = 36.1
Nominal : [ 1104.1, 1104.7, 640, 360 ]